Page 1

| With thanks to: McMaster University Dept. Pathology and Molecular Medicine Center for Gene Therapeutics Standard Operation Procedure Ultracentrifuge |
| --- |
| Initial Issue date: May 5 /2005 Revision Date: Oct.12/ 2007 |  | Section: Center for Gene Therapeutics |
| Approved by Jack Gauldie: (Supervisor) Carol Lavery(?): (CGT-Lab Manager) | Page 1 of 5 |

##### PURPOSE

This procedure describes the stepwise standard operation for using an ULTRACENTRIFUGE.

## RATIONALE

The ultracentrifuge is a delicate piece of equipment. Rotors on ultracentrifuge units are subject to mechanical stress that can result in rotor failure. In addition, improper loading and balancing of the rotor can cause it to break loose while spinning. For these reasons, the centrifuge must be properly used and maintained.

Failure to follow SOP may not only result in the damage to the equipment and biological samples, but also cause the injury to the person who is operating it. To get familiar with the function, specifications, operation and routine operator care and maintenance is essential to the proper use of the ultracentrifuge and the safety of people.

##### SCOPE

All trained technicians, research fellows and students are required and responsible to follow the SOP for operating and maintaining the ultracentrifuge.

##### REQUIRED FORMS

Monthly calendar sheets for RESERVE booking of centrifuge and rotor.

Rotor log book to record run information.

The Beckman CoulterTM Preparative Ultracentrifuge Log Sheet. (See Appendix 1)??

## REFERENCES

Beckman Coulter Optima L-90K Preparative Ultracentrifuge Instruction Manual.

## Beckman Coulter Ultracentrifuge Rotor Instruction Manuals.

## EQUIPMENT

Beckman Coulter OptimaTM  L-90K Preparative Ultracentrifuge

Beckman    SW 50.1   Rotor   S/N    96U6250    P/N   340079

Beckman    SW 41 Ti  Rotor   S/N    97U9939    P/N   331362

Beckman    SW32Ti    Rotor   S/N    02U 441  P/N   369650

## PROCEDURE

PRE-START

BOOK CENTRIFUGE AND ROTOR ON CALENDAR PROVIDED

PRECHECKING

Check that the centrifuge chamber, drive spindle, and tapered mounting surface of the rotor are clean and free of scratches or burrs.

Examine the centrifuge tubes and buckets for cracks and mechanical defects. Wipe the drive surfaces prior to installing the rotor. Apply a “very small” amount of Spincote to spindle thread and metal on metal contact points. See instruction manual for the rotor you are using. Check and replace any worn or damaged o-rings. Apply a “small amount” of  BECKMAN vacuum grease to o-rings to ensure a proper seal.

Do not exceed the maximum rated speed of the rotor in use.

STANDARD OPERATION

1, Balance the loaded tubes carefully with electronic scale in biological safety cabinet and make

sure the tubes are well balanced. Avoid over filling or under filling.

2, Load tubes into the buckets, place bucket caps on and screw caps tight with screwdriver.

Transport on the rail cart to the centrifuge room. RUN ROTORS WITH A FULL SET OF

BUCKETS.

3, Turn on the power, open the chamber door, install the rotor and close the chamber. Press

VACUUM if you want to evacuate the chamber.

4, Enter the required run conditions: SPEED, TEMP, and TIME, HOLD, or (2t.

5, Select ACCEL and DECEL profiles if required, otherwise, the instrument will

automatically select maximum acceleration and deceleration rates.

6, Press ENTER/RECALL then START. (The vacuum system is automatically activated if you

did not press VACUUM earlier.

7, Press STOP to terminate a run in the HOLD mode (or to stop any run in progress.)

Runs in the timed or (2t mode will terminate automatically when the set value is reached.

8, When the rotor has stopped, press VACUUM to vent the chamber. Then open the door and

remove the rotor.

9, Leave the door open to evaporate any water after refrigerated runs and switch the power off.

10, Record the run information in log book for the rotor/ serial # used. Failure to comply will       invalidate Beckman rotor warranties.

###### REPEAT THE SAME RUN

To repeat the same run, press ENTER/RECALL then START. There is no need to re-enter run

condition unless you need to make a change.

###### PROGRAMMED OPERATION

1, Press PROG, then use the keypad to select program number. (Or press PROG until the

required program appears.)

2, Press ENTER/RECALL then START. (The vacuum system is automatically activated if you

did not press VACUUM earlier.)

To Change a Parameter During a Programmed Run

1, Press the appropriate parameter key (for example, SPEED).

2, Use the keyboard to enter a new value.

3, Press ENTER/RECALL. Change other value in the same manner as required. The run in

progress will be modified accordingly.

4, Run in the time or (2t mode will terminate automatically when the set value is reached.

When the rotor has stopped, press VACUUM to vent the chamber. Then open the door and

remove the rotor.

DELAYED START PROGRAM

1, Press PROG. Use the keypad to enter 0 and then the number of the second program. Press

ENTER/RECALL.

2, Press ENTER/RECALL then start.

CLEAN-UP PROCEDURE

After use, buckets, rotors and centrifuge interiors should be cleaned and/or disinfected. Clean up any biologic spills immediately. Follow appropriate spill response procedures. Don’t use corrosive reagents such as strong acid or base and chlorides to clean the buckets. Use the Beckman rotor cleaning detergent provided to clean rotors and buckets. Rinse buckets with 70% ethanol; wipe large buckets dry with kimwipes; invert small buckets on absorbent material and replace lids when dry.

Page 4

TROUBLE SHOOTING

Review the owner’s manual to identify and solve the problem that appears to you. Contact and report to the responsible person if you cannot solve the problem. Never attempt to repair yourself. Centrifuge should be repaired only by the manufacturer or authorized dealer representatives.

DOCUMENTATION

Record of using and maintaining the Ultracentrifuge is

reviewed and checked on a regular basis by quality control personnel, and retained in the

laboratory.

PREVENTATIVE MAINTENANCE

Ultracentrifuges are under contract with Beckman for 2 PM visits per year and 24 hour emergency service.

Rotors are inspected for wear or damage by Beckman service annually and de-rated or decommissioned as appropriate. Reports are maintained in the lab.

It is the responsibility of the USER to report any damage or malfunction of this equipment to the emergency contact person listed.  You are required to leave a note on the equipment describing the problem and including your name and date the problem occurred.

Appendix 1

#### PREPARATIVE ULTRACENTRIFUGE LOG SHEET

| Date | Revolution Counter At End of Run | Rotor Type | Rotor Serial NO. | RMP | Run Length | Operator | REMARKS (oil level, temperature, density gradient, etc.) |
| --- | --- | --- | --- | --- | --- | --- | --- |

# Text boxes

Page 2

Page 3